## Sheet1
FROM EQUIP LIST: | NEW BLDG | GENERAL | [5 empty] | DIMENSIONS |  |  |  | FLOOR OR BENCHTOP | CLEARANCE |  |  |  | ELECTRICAL | [7 empty] | MECHANICAL | [5 empty] | PLUMBING |  |  |  |  | STRUCTURAL | VIBRATION REQUIREMENTS | NOISE (dbs) | CLEAN ROOM REQUIREMENTS |  |  |  | REMARKS/ ADJACENCIES
|  | USER GROUP NAME | ID | ASSET TAG | ITEM NAME | MANUFACTURER | MODEL NUMBER/ DESCRIPTION | WEIGHT (LBS) | WIDTH IN INCHES | DEPTH IN INCHES | HEIGHT IN INCHES |  | FRONT IN INCHES | BACK IN INCHES | SIDE IN INCHES | MAINTENANCE RELIEF | RECEPTACLE OR PLUG TYPE | VOLTAGE (RECEPTACLE OR EQUIPMENT) | PHASE | AMP | WATT OR VA | CAPACITY FACTOR | HERTZ | ELECTRICAL UTILITIES list from key | SUPPLY CFM | SUPPLY PSI | EXHAUST CFM | HEAT BTU/hour | MECHANICAL UTILITIES list from key | AIR | DRAIN | WASTE | PLUMBING UTILITIES list from key | MAX PSI | GPM |  |  |  |  | RELATIVE HUMIDITY | TEMPERATURE | CLEANLINESS CLASSIFICATION
GEN | 100-1 | LAB | FA-1 | B-1 | EQ1 | GE | tbd | 18078 | 100" 140" clearance for R-hand door swing (dims in cut sheet and tech data doc match) | 80" to console front 100" total with generator cabinet (used tech data dims - bigger) | 80* (dims match in all docs) | FLOOR | [5 empty] | 230** **comes with isolation transformer - step the building power up from 208 to 230 | 1 | 9 |  |  |  | DA | 0.88 | 90*** | [10 empty] | vibration sensitive | 75 |  |  |  |  | 3D reconstructed XRAYs
GEN | 100-2 | LAB | FA-1A | B-1A | EQ2 | GE | tbd | 771 (from tech data - higher than cut sheet) |  |  |  | FLOOR | [24 empty] | minimal to none created
GEN | 100-3 | LAB | FA-1B | B-1B | EQ3 | GE | tbd |  |  |  |  | FLOOR
GEN | 100-4 | LAB | FA-1C | B-1C | EQ4 | GE | tbd | tbd | Standard desk - 60" x 36" |  |  | FLOOR | [5 empty] | 120V wall outlet for computer | [5 empty] | DA
GEN | 100-5 | LAB | FA-2 |  | EQ5 | GE | tbd |  | 120 | 144 |  | FLOOR |  | 48 | 60 |  |  | 110 |  |  |  |  | 60 | [14 empty] | < 45 dB |  | < 75% | AC 20-25C +/-1C |  | UPS ~2 KVA, Laser safety signs